1. Title and Approval

- Your Specific Project Under the “Vermont Agricultural and Environmental Laboratory (VAEL) Services Grants Project - 2016”:

_______________________________________________________________________

(Your Project’s Name)

________________________________________________________________________

(Name of Your Organization)

_______________________________________________________________________

(Date)

Project Manager Signature/Date:

Project QA Officer Signature/Date:

Project QAPP Prepared by:

Approval by:

James Kellogg

VTDEC LaRosa Analytical Services Grant Partnership Coordinator

Date

- Generic Volunteer-Based Water Quality Monitoring Project QAPP:

Vermont General Quality Assurance Project Plan for Volunteer, Educational and Local Community Monitoring and Reporting Activities

(Project Name)

VT Department of Environmental Conservation

(Responsible Agency)

April 1, 2016

(Date)

QAPP Prepared by: Lee Steppacher & Diane Switzer, EPA New England, Modified by James Kellogg, VTDEC.

2.  Table of Contents

| Section |  | Page |
| --- | --- | --- |
| 1. | Title and Approval Pages The Specific Project Generic Volunteer Based Water Quality Monitoring QAPP | 1 2 |
| 2. | Table of Contents | 3 |
| 3. | Distribution List | 5 |
| 4. | Project/Task Organization A. VTDEC - Key People and Responsibilities Project Key People and Responsibilities | 6 |
| 5. | Background of Project | 7 |
| 6. | Individual Project Purpose/Task Description Objectives of Projects Intended Uses of Data Map of Area and Waterbody Project Timetable | 8 |
| 7. | Project Quality Objectives Data Precision, Accuracy, Measurement Range Requirements Data representativeness Data Comparability Data Completeness | 12 |
| 8. | Training Requirements and Certification | 16 |
| 9. | Documentation and Records | 17 |
| 10. | Sampling Process Design Rationale for Selection of Sampling Sites Summary of Sampling Collection | 17 |
| 11. | Sampling & Analysis Methods. | 23 |
| 12. | Sample Handling and Custody Procedures | 25 |
| 13. | Analytical Methods Requirements | 25 |
| 14. | Quality Control Requirements Field QC Checks Laboratory QC Checks | 25-26 |
| 15. | Instrument/Equipment Testing, Inspection and Maintenance Requirements | 26 |
| 16. | Instrument Calibration and Frequency | 28 |
| 17. | Inspection/Acceptance Requirements | 28 |
| 18. | Data Acquisition Requirements | 29 |
| 19. | Data Management | 29 |
| 20. | Assessment and Response Actions | 29 |
| 21. | Reports | 30 |
| 22. | Data Review, Validation and Verification | 30 |
| 23. | Validation and Verification | 31 |
| 24. | Reconciliation with Project Quality Objectives | 31 |
| Appendices | (Hard copies of appendices are available from the VTDEC) |  |
| A. | Examples of Acceptable Field Standard Operating Procedures (Field SOPs) |  |
| B. | Examples of Acceptable Field Sheets and LaRosa Laboratory Sample Submission Form |  |
| C. | VTDEC Citizen’s Guide to Bacteria Monitoring |  |

3.  Distribution List

- Names and telephone numbers of those receiving copies of this QAPP.

- Jim Kellogg, VT Department of Environmental Conservation.  Watershed Management Division, 1 National Life Drive, Main Building - 2cd Floor, Montpelier, VT 05602-3522. 1(802) 490-6146. Jim.Kellogg@vermont.gov

- _______________________________________________________

- ________________________________________________________

- _______________________________________________________             _

- ____________________________________________________                __

4.  Project/Task Organization

- Table 4a – VTDEC/VAEL – Primary Contact and Their Responsibilities.

| Project Title/Responsibility | Name |
| --- | --- |
| Project Coordinator for VTDEC | Jim Kellogg |
| Volunteer Monitoring Coordinator for VTDEC - Watershed Management Division (WsMD) | Bethany Sargent |
| VAEL Director | Guy Roberts |
| VAEL Supervisor | Dan Needham |

- Table 4b - Key Project People and Their Responsibilities

| Project Title/Responsibility | Name/Affiliation |
| --- | --- |
| Project Leader – directly responsible for all aspects of this project. |  |
| Project Volunteer Coordinator – responsible for overseeing all of the activities of volunteers, including recruiting, maintaining training and participation records, etc. |  |
| Project Field/Sampling Leader – responsible for training and supervising volunteers in field work, filling out field forms, and performing QC checks to make sure procedures are followed or corrected, as needed. |  |
| Project QA Coordinator – responsible for ensuring that procedures in field and laboratory are performed in accordance with this QAPP and keeps other leaders informed of project status in relation to QAPP. Works with other leaders in conducting QC checks on sampling and analysis techniques. Maintains log of QA and QC checks. |  |
| Project Laboratory Contact – makes arrangements with laboratory’s used to perform analysis according to QAPP. Ensures correct procedures are used, holding times are met and laboratory provides complete documentation. |  |
| Project Data Management Coordinator – Maintains the data systems for the organization, performs data entry, and checks entries for accuracy against field and laboratory forms. |  |

5.  Background of the LaRosa Analytical Services Grant Partnership

The Vermont Department of Environmental Conservation (VTDEC), through the Vermont Agriculture and Environment Laboratory, has made available to interested lake, river, and watershed associations grants for sample analyses since the 2003 field season.  The purpose of this program is to help watershed associations and volunteer monitoring groups to implement new and/or on-going surface water monitoring projects, for waters in need of water quality assessment.

What are laboratory services?

One of the costliest items involved in a monitoring program is laboratory analysis.  VTDEC recognizes that the cost of laboratory services hinders the widespread application of volunteer surface water quality monitoring in Vermont.  Analytical services provided under this grant program are essentially ‘slots’ for tests to be run at VAEL, free of charge to LaRosa grant recipients.  VAEL is a full-service analytical facility with complete capabilities for routine water quality monitoring tests.  Examples of such tests include: phosphorus, nitrogen, chlorophyll-a, total suspended solids, E. coli, turbidity, alkalinity, conductivity, pH, priority pollutants metals and numerous other compounds.  More information about the VAEL’s services are available online (http://www.anr.state.vt.us/dec/lab/index.htm).

Who is eligible?

Volunteer associations across Vermont are eligible for this project.  Such associations include river, lake, and watershed groups, secondary-level educational groups, and water quality and conservation committees associated with local municipalities.  Post-secondary academic institutions and statewide not-for-profit non-governmental organizations are eligible provided that the projects are either: designed jointly with a local association to assess current water quality conditions; or, structured to address a water quality problem of statewide importance.

What are the eligible project types?

Many project types are eligible for this program.  Waters under evaluation should be of interest to the local association sponsoring the project, and should also be of interest to VTDEC.  Waters of interest to VTDEC include impaired and state priority waters, waters on which minimal or no monitoring has been performed in the past, waters with significant public swimming use, waters where a suspected water quality problem needs further assessment, and waters where the causes of known problems remain undiagnosed. Proposals for projects exceeding one field season in duration will be accepted, although subsequent years will be approved only subject to continued availability of state funding for this program.  Please note that participants in this program shall share with VTDEC ownership of all laboratory data produced by individual projects.

- Individual Project Purpose/Task Description

Instructions - For Parts A and B below, please check the boxes that apply to your project and add specific information as needed. Include all pertinent background information that helps support the purpose of your project, including a brief summary of previously collected data.  The summary can either be in table format or a brief narrative.

Attach a map in Part C, to identify waterbodies being sampled and sampling sites.  If you are unable to locate sampling sites until the project is initiated, please explain your circumstances below.

- Objectives of Projects

The principal objectives of projects under this QAPP are to 1) provide a perspective on the range of water quality conditions across Vermont; 2) describe water quality conditions of individual waterbodies; 3) establish a data base for waterbodies for use in documenting future changes in water quality; and, 4) educate and involve local residents in waterbody protection.

General guidelines for projects under this QAPP are:

- Data should be collected during summer months at regular intervals, but not in severe weather, such as thunderstorms or high winds (safety always comes first).  Projects addressing E. coli should be designed specifically to address either dry-only weather conditions, or segregate between wet and dry weather conditions.  Antecedent weather conditions must be recorded for all E. coli sampling events.

- If some data will be collected every week, and other data will be collected only once during the sample season or appropriate index period (e.g., low flow, high temperature, etc.), such should be noted in Section 10B, Sample Design Logistics, in this QAPP.

- Data will be analyzed, summarized and interpreted on an annual basis. Projects will be required to report to VTDEC at the completion of the project.

- Information should be presented to the local community in a suitable format, be it a press release, public meeting, or another event.

- Data that meets project quality objectives may be entered into VTDEC’s Water Quality data management system and may also be entered into EPA’s national water quality data storage system, STORET.

- Intended Uses of Data

Instructions: Please place a checkmark beside the uses which are applicable for your project’s data.

The data generated by projects under this Generic QAPP will serve at least one of the following uses, as specified in project proposals and work plans.

- Track phosphorus concentrations and/or loadings

- Identify the presence, density and spread of nuisance aquatic species

- Describe water quality conditions at specific locations

- Document the presence and severity of localized problems (e.g. bacteria as pathogen indicators)

- Identify sources of local problems

- Evaluate sedimentation and erosion problems

- Evaluate habitat & embeddedness with regard to aquatic life use

- Educate school children and local communities about water quality, and any problems and improvements.

- Evaluate the effectiveness of restoration projects and other management activities

- _______________________________________________________________

________________________________________________________________

________________________________________________________________

- Map of Area and Waterbody

For individual projects under this generic QAPP, a map is to be provided here that identifies waterbody and sample sites.

Instructions: Insert the map for your project here.

- Table 6a - Project Timetable

| Activity | Projected Start Date | Anticipated Completion Date |
| --- | --- | --- |
| Project Planning Meeting |  |  |
| Fill out and submit this QAPP to VTDEC |  |  |
| QAPP Approved by VTDEC |  |  |
| Training Volunteers/Samplers |  |  |
| Sampling Begins |  |  |
| Sampling Ends |  |  |
| Analytical Results Evaluated * Check/Correct Errors Due to Math Miscalculations or Transferring Data from Field/Lab Forms * Confirm Useable Data * Separate Unusable Data |  |  |
| Data Entered into Project Database |  |  |
| QC Review of Database |  |  |
| Data Summarized |  |  |
| Submit Final Report |  |  |
| Presentation(s) of Information at Local Meeting (s) or other venue(s) |  |  |

7.  Project Quality Objectives

Instructions: Please check to ensure that you can meet the accuracy and precision requirements, and if you cannot please indicate and explain. Check the appropriate boxes on the left for parameters to be sampled in your project.  If you plan to use a different field or laboratory method add your information to this table and provide the written procedures when submitting this completed project QAPP.

A.  Data Precision, Accuracy, Measurement Range Requirements

Table 7a –   Field Analysis Protocols for Water Samples

| Parameter | Field Analysis Method | Method Reference¹ | Accuracy2 | Precision2 |
| --- | --- | --- | --- | --- |
| Transparency | Secchi Disk | Vermont Lay Monitoring Program Manual, 2000 | -- | +/- 0.1 meter |
| Dissolved Oxygen by Meter | DO Meter or multiprobe | Standard Methods for the Examination of Water and Wastewater, 20 ed., 4500-O G. Membrane Electrode Method | +/- 0.5 mg/l | +/- 0.5 |
| Temperature | Alcohol Thermometer | Testing the Waters; Chemical & Physical Signs of a River, River Network,1997 | +/- 1.0º C | +/- 1.0º C |
| pH | pH Meter or multiprobe | Standard Methods, etc.,20th ed., 4500-H+B Electrode Method | ± 0.2 std.un. | ± 0.2 std.un. |

Footnotes:

1–The full citations for each of these publications are:

APHA, AWWA & WEF.  Standard Methods for the Examination of Water and Wastewater, prepared and published jointly by the American Public Health Association, American Water Works Association and Water Environment Federation, 20th ed., 1998

Behar, Sharon.  Testing the Waters; Chemical & Physical Vital Signs of a River, published by River Network, 1997

Vermont Agency of Natural Resources.  Vermont Lay Monitoring Program Manual; 2000, by Water Quality Division, Vermont Dept. of Env. Conservation.

2–  Accuracy of field protocols will generally not be measured in the field, but at training and quality control check sessions.  Accuracy and Precision measures given are generic.  Individual protocols may themselves provide more accurate and precise measures than expressed here.

Table 7b –  Laboratory Analysis Protocols for Water Samples:

| Parameter | Reporting Limit A | AccuracyB (% Recovery) | Estimated Precision for Field Duplicates C (RPD) | Laboratory Precision (RPD) | Analytical Method ReferenceB |
| --- | --- | --- | --- | --- | --- |
| Chlorophyll-a | 0.5 ug/l | -- | ≤15% | 10% | EPA 445.0 |
| Total and dissolved phosphorus | 5 (g/l | 85-115% | ≤30% | 15% B | Std. Methods (21st ed.) 4500-P H |
| E. coli D, E | 1 MPN /100ml | N/A | 125% (<25cfu) 50% (>25 mpn) | 125% (<25cfu) 75% (>25 mpn) | Std. Methods (21st ed.) 9223 (Colilert) |
| Chloride | 2 mg/l | 85-110% | ≤ 5% | ≤ 5% | Std. Methods (21st ed.) 4500-Cl G |
| Total Suspended Solids | 1 mg/l | 80-120% | ≤15% | ≤ 15% | Std. Methods (21st ed.) 2540D |
| Turbidity | 0.2 NTU | N/A | ≤ 15% | ≤15% | EPA 180.1 |
| Alkalinity | 1 mg/l | N/A | ≤5% (>20 mg/l)<15% (<20 mg/l) | ≤5% (>20 mg/l)<15% (<20 mg/l) | Std. Methods (21st ed.) 2320B |
| Total nitrogen (persulfate digestion) | 0.1 mg/l | 85%-115% | ≤20% | ≤10% | Std. Methods (21st ed.) 4500-N C |
| Total NOx | 0.05 mg/l | 85%-110% | ≤10% | ≤5% | EPA 353.2 |

(A) - Reporting Limit is the minimum reported value (lowest standard in calibration curve or MDLx3)

(B) - Section 5.0, Vermont Dept. of Conservation Laboratory QA Plan, 2008

(C) - Generated by the analysis of field duplicates

(D) - EPA’s New England Regional Laboratory recommends that all samples resulting in Too Numerous to Count (TNTC) growth, defined as greater than 200 colonies on the membrane filter, be recorded as “TNTC.”

(E) -As a quality control check on bacteria counts, if two or more analysts are available, each should count colonies on the same membrane plate for about 10% of the samples, and agree on the # of colonies within 10%.

B.  Data Representativeness

Samples collected at locations and depths described in this QAPP will reflect conditions of individual waterbodies and tributaries in Vermont. To ensure representativeness all samples will be collected, preserved and analyzed according to the procedures in this QAPP, and within the specified holding times.   Those results not meeting the project quality objectives of this program will be flagged and reviewed to determine if appropriate quality controls are in place.  They should be discussed in the data report and may be excluded from entry into VTDEC’s long-term water quality data archive referred to as WQX.

C.  Data Comparability

NOTE: The information in Table 7c – Project Completeness (below) about field samples, and field and lab duplicate samples collected, is not needed for the QAPP submission; however, please review it so you will be able to submit it at the end of the project.

All samples for each specific parameter will be collected and analyzed using the respective procedures described in this QAPP to ensure that comparisons between different sample sites, sample dates, depths and projects can be appropriately made.

If a project compares historical data with the data generated under this QAPP, the historical data should have used SOPs that provide the same data quality as defined here.

D.  Data Completeness

At least 80% of the anticipated number of samples will be collected, analyzed and determined to meet data quality objectives for the project to be considered successful.  Individual projects may have different completeness goals, which will be presented in the table below.  The data report for each project will contain information, similar to that presented below, containing the number of samples meeting the data quality objectives and the resulting calculation of “Percent Complete”.

__________________________________________________________________________________________________________________________________________

Table 7c – Project Completeness

| Parameter | Number of Samples Anticipated | Number of Valid Samples Collected & Analyzed | Percent Complete * |
| --- | --- | --- | --- |
| Chlorophyll-a |  |  |  |
| Chloride |  |  |  |
| Total and Dissolved Phosphorus |  |  |  |
| E. coli |  |  |  |
| Total Suspended Solids |  |  |  |
| Transparency |  |  |  |
| Alkalinity |  |  |  |
| pH |  |  |  |
| Turbidity |  |  |  |
| Total nitrogen (persulfate digestion) |  |  |  |
| Total NOx |  |  |  |
| Si, dissolved |  |  |  |
| Dissolved Oxygen |  |  |  |
| Conductivity |  |  |  |
| Temperature |  |  |  |

*   Percent Complete = # of Valid Samples Collected and Analyzed / # of Samples Anticipated

8.  Training Requirements and Certification

- Training Logistical Arrangements

Instructions: Make changes as needed to the table below to reflect your project.  Note however that what is contained in this table is, for the most part, considered minimal training.

The Project Leader will arrange volunteer training sessions and keep a record of each volunteer’s training needs and accomplishments. Project Leaders are encouraged to discuss their training needs with the VTDEC-WQD Volunteer Monitoring Coordinator.

Table 8a - Training Process

| Type of Volunteer Training | Frequency of Training/Certification |
| --- | --- |
| Initial Orientation to the Project | Once |
| Training in Sampling and Analysis | One full training session for each new volunteer & annual refresher training before each sampling season begins |
| On-site Visit by Project Coordinator | Once during sampling season |
| Other: |  |

9.  Documentation and Records

Documentation for each project will include 1) sample forms 2) field sheets and 3) written assessments from on-site visits of Project Leader & QA Coordinator (see Section 8A).  The Project Leader will maintain a record of each volunteer’s training and participation in projects.  Field data sheets will be filled out by the Sampling Volunteer and maintained by the Project Manager.  Example field sheets are included in Appendix B.  Each group will attach a copy of their field sheets to the individual project’s version of this QAPP they submit.  All samples submitted for laboratory analysis will be accompanied by VTDEC’s sample submission form (Appendix B).

10.  Sampling Process Design

- Rationale for Selection of Sampling Sites

SAMPLE SITE DESCRIPTION – Please provide a general description of each sample site, and note the approximate location on the submitted map .

| Site # | Description |
| --- | --- |

PHYSICAL HABITAT & SURROUNDING FEATURES – Characteristics of the physical habitat, land use in the immediate area, or specific features like point source discharges, may help determine where sample sites are located (e.g., macroinvertebrate sampling may take place only in riffle areas). Where this is the case, please describe the rationale for site selection.  The intensity of the description will depend on individual projects, and must meet the requirements necessary to use the data for the project’s purpose.

| Site # | Habitat/Surrounding Features |
| --- | --- |

LOCATIONAL DATA – The latitude/longitude of each sample site will be recorded using a Global Positioning System.  If this is not available, map coordinates including the map datum from which the coordinates were derived must be provided.  This information does not need to be submitted as part of the QAPP, but the information must be included in the project data report.

| Site # | DD.mmmmm North | DD.mmmmm West |
| --- | --- | --- |

LAKES/PONDS WATER QUALITY–   Generally, each Lake/Pond will be sampled for water quality parameters at the deepest spot.  For some projects, sampling will take place at the center of the lake, regardless of whether it is the deep hole, depending on the specific purpose of the monitoring event and the parameters being assessed.  If sampling for your project will vary from any of these designs, please describe it below.

RIVERS/STREAM WATER QUALITY – Wadeable stream samples will generally be collected offshore, near the center of the stream.  Water quality samples will be taken from just below surface to near bottom.  Individual grab samples, composited grab samples or a core sample can be collected from the water column.  Specific projects will designate the type of sample, which must be in accordance with quality control requirements and the purpose of each project.

Depending on the bottom substrate, water quality samples from deep rivers should be collected at mid-depth, but no closer than 0.5 meters from the sediment interface.  If the substrate is very soft/silty a greater distance may be designated so as not to contaminate the water sample or the sampling device

Instructions: Please check the types of samples that will be collected for your project.  If your sampling method(s) differ from the description, please describe what you intend to do.

For this specific project, the samples will be collected by:

- Individual grab samples that will be analyzed separately

- Time composite samples – the same volume is collected at constant time intervals (e.g., 4 hours apart) at the same site, and combined to form a composite sample for that site

- Core samples – a single sample collected vertically in the water column across a series of depths.

If sampling for your project will vary from this design, please describe it below.

______________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

Table 10a – Overview of Types of Waterbody, Sample Site(s) & Sample Depth(s)

| TYPE OF WATERBODY | SAMPLE SITE(S) For Each Waterbody | SAMPLE DEPTH(S) At Each Site | TRANSECT(S) Across Length or Width of Each Waterbody |
| --- | --- | --- | --- |
| How many LAKES & PONDS will be sampled? ____ |  |  |  |
| Name of Lake/Pond: | Inflow Outflow Center Deep Hole Near Shore | Surface Water Mid-Depth Near Bottom Bottom Water Surface to Bottom Profile Bottom Substrate | Transect from Inflow to Outflow Cross Transect(s) Transect(s) near to shore & parallel to shoreline |
| Name of Lake/Pond: | Inflow Outflow Center Deep Hole Near Shore | Surface Water Mid-Depth Near Bottom Bottom Water Surface to Bottom Profile Bottom Substrate | Transect from Inflow to Outflow Cross Transect(s) Short Transects parallel to shoreline Short transects extending out from shoreline |
| How many RIVERS & STREAMS will be sampled? _______ |  |  |  |
| Name of River/Stream: | Upstream of __________ Downstream of ___________ Wadeable Deepwater | Surface Mid-Depth Near Bottom Bottom Surface to Bottom Profiles Bottom Substrate | Upstream to Downstream Transect Cross Transect Name of River/Stream: |
| Name of River/Stream: | Upstream of __________ Downstream of ___________ Wadeable Deepwater | Surface Mid-Depth Near Bottom Bottom Surface to Bottom Profiles Bottom Substrate | Upstream to Downstream Transect Cross Transect |
| Name of River/Stream: | Upstream of __________ Downstream of ___________ Wadeable Deepwater | Surface Mid-Depth Near Bottom Bottom Surface to Bottom Profiles Bottom Substrate | Upstream to Downstream Transect Cross Transect |
| Name of River/Stream: | Upstream of __________ Downstream of ___________ Wadeable Deepwater | Surface Mid-Depth Near Bottom Bottom Surface to Bottom Profiles Bottom Substrate | Upstream to Downstream Transect Cross Transect |
| Name of River/Stream: | Upstream of __________ Downstream of ___________ Wadeable Deepwater | Surface Mid-Depth Near Bottom Bottom Surface to Bottom Profiles Bottom Substrate | Upstream to Downstream Transect Cross Transect |

- Summary of Sample Collection

Individual projects will identify the number of samples, sampling frequency and specific sampling method for each parameter in accordance with their objectives.  During sample collection, all sample apparatuses are to be rinsed 3x in sample water prior to collection of the actual sample (except where noted).  Filtration apparatuses and bottle rinse guidelines are shown in Table 11a.

Table 10b – Sample Collection

|  | Type of Sample/ Parameter | Total Number of Samples (Indicate if this is for the project or per week, etc.) | Sampling Frequency (How often – once/weekly/bi-weekly?) | Sampling Method (Grab, Discrete-depth sampler, depth-integrating core sampler, meter) * |
| --- | --- | --- | --- | --- |
| Biological | E. coli |  |  |  |
| Chemical | Chlorophyll-a |  |  |  |
|  | Chloride |  |  |  |
|  | Total and Dissolved Phosphorus |  |  |  |
|  | Transparency |  |  |  |
|  | Dissolved Oxygen |  |  |  |
|  | Temperature |  |  |  |
|  | pH |  |  |  |
|  | Alkalinity |  |  |  |
|  | Total Nitrogen (persulfate digestion) |  |  |  |
|  | Total NOx |  |  |  |
|  | Si, dissolved |  |  |  |
| Physical | Secchi Disk Transparency |  |  |  |
|  | Total Suspended Solids |  |  |  |
|  | Turbidity |  |  |  |
| Meters used for data collection (please list make/model of meter(s) or multiprobe(s) | Multiprobe model: |
|  | pH meter model: |
|  | Conductivity meter model: |
|  | Turbidity meter model: |
|  | DO meter model: |

* see Appendix A, please list sampler type (e.g., Kemmerer, Van Dorn, Hose etc.).

11.  Sampling & Analysis Methods

Field and laboratory analytical methods are provided in Section 7, and Field Sampling Methods are listed in Section 10 and in Appendix A.  The table below presents containers, preservation and holding times used for projects under this QAPP.

Table 11a –Sample Containers, Preservation & Holding Times A

| Parameter/Measure | Container | Field Rinse | Preservation | Hold Time B |
| --- | --- | --- | --- | --- |
| q Total / Dissolved Phosphorus | 60 ml glass tubeC | NO RINSE, 3X rinse of filtration apparatus w/ sample water or DI | Dissolved phosphorus filtered using new 0.45 u filter membrane | 28 days |
| q E. coli | 290ml or 120ml sterile plastic round | NO RINSE | Cool to <10°C | 8 hours |
| q Chlorophyll-a | Filter - Whatman GF-F, 47mm diam., 0.7 µm pore size, stored in black jar | NO RINSE of filter, 3X Rinse of filtration apparatus w/ sample water or DI | Freeze (20 to -70ºC), Dark | 21 days |
| Chloride | 50 ml polycarbonate centrifuge tube | 3x rinse with sample | Cool to <6°C | 28 days |
| q Total Suspended Solids | 1L plastic, round | 3x rinse with sample | Cool to <6°C | 7 days |
| q Turbidity | 250 ml plastic square | 3x rinse with sample | Cool to <6°C | 48 hours |
| q Total Nitrogen (persulfate digestion) | 50 ml polycarbonate centrifuge tube | 3x rinse with sample | Cool to <6°C, acidified within 48h with conc. H2SO4 to pH <2 | 28 days |
| q Total NOx | 50 ml polycarbonate centrifuge tube | 3x rinse with sample | Cool to <6°C, acidified within 24h with conc. H2SO4 to pH <2 | 28 days |
| q Si, dissolved | 50 ml polycarbonate centrifuge tube | 3x rinse with filtrate or with DI | Cool to <6°C, filter using new 0.45um filter membrane | 28 days |
| q Alkalinity | 250 ml plastic square | 3x rinse with sample | Cool to <6°C | 14 days |
| q DO - Meter | (in situ) | 3x rinse of probe | None | Direct Analysis |
| q pH Meter | (in situ) | 3x rinse of probe | None | Direct Analysis |
| q Temperature - Thermometer D or meter | (in situ) | NO RINSE | None | Direct Analysis |
| q Conductivity meter | (in situ) | 3x rinse of probe | None | Direct Analysis |
| q Turbidity meter | (in situ) | 3x rinse of probe | None | Direct Analysis |

Footnotes:

A – A copy of some field SOPs are attached as Appendix A.

B – Holding times are in accordance with the Code of Federal Regulations, title 40 (Protection of Environment), part 136, section 3 (or 40CFR136.3), and are defined in the VTDEC LaRosa Laboratory Quality Assurance Project Plan.

C – The VT DEC analyzes the entire sample volume in the sampling container, so no acidification is needed.  Extra containers of sample will be needed to allow the VT DEC lab to analyze spiked samples.

D – Mercury thermometers absolutely shall not be used in the field.

12.  Sample Handling and Custody Procedures

All samples collected in conjunction with the project will be accompanied by a field sampling form identifying at minimum the sample location, date, time, and collector.  In addition, a laboratory sample submission form must accompany all samples submitted to the laboratory. Sample field forms, and a laboratory submission form, are in Appendix B.

13.  Analytical Methods Requirements

Information for this section is included in Tables 7a and 7b.

14.  Quality Control Requirements

A.  Field QC Checks

At least one Field Duplicate and one Field Blank will be submitted for every ten samples collected.  Additional types of field quality control samples needed will depend on the parameter and the collection method, and are at the discretion of the Project Manager and QA Manager.

- Field Duplicate (required) – a check on water quality, sampling & analysis consistency.  This is a replicated sample collected at the same point in time and space so as to be considered identical. A field duplicate is a second sample from a second sampling event, collected immediately after the first sampling. Otherwise put, these separate samples are said to represent the same population and are carried through all steps of the sampling and analytical procedures in an identical manner.  They are used to assess precision of the total method, including sampling, analysis, and site heterogeneity.

- Field Blanks (required) – This checks for contamination (Accuracy/Bias) in the field by processing laboratory-supplied deionized through the sampling train.  This checks for contamination introduced from the sample container(s) or from field contamination.

- Matrix Spike (required only for phosphorus) - This allows the laboratory to perform analytical replication that separates variability in sampling from variability in analytical processing.  A spike is a second sample bottle, filled from the same sample collection as the first sample.  For grab samples, there is no functional difference between a field duplicate and a matrix spike.

- Equipment Blanks – measures contamination (accuracy/bias) – a sample of water, free of measurable contaminants, is poured over or through decontaminated field sampling equipment that is considered ready to collect or process an additional sample.  The purpose of this is to assess the adequacy of the decontamination process and whether equipment needs special cleaning to make sure it doesn’t have something that contaminates the sample or influence the results

- Field Split Samples – Two or more representative subsamples are taken from one environmental sample in the field and sent to two different labs for analysis.  Prior to splitting, the environmental sample is well-mixed to correct for sample inhomogeneity that would adversely impact sample data comparability. Field splits are used to assess sample handling procedures from field to laboratory and inter-laboratory comparability and precision.

- Equipment Calibration Checks – A check on a meter’s accuracy – the verification of the initial calibration that is required at certain times during the sampling day or while analyzing a large number of samples.  Checking to see if a pH meter is maintaining its calibration would involve taking a reading of standard solutions (e.g., pH buffers of 4, 7, or 10, etc.)

For projects that include long term repetitive sampling at several sites, the site at which a field quality control sample is collected may change so there may be at least one duplicate sample at each sample location during the course of the project.

B. Laboratory QC Checks

Laboratory QC samples may include any of the following, depending on the parameter, and are handled by the VAEL as described in the VAEL Quality Assurance Plan.

15.  Instrument/Equipment Testing, Inspection, and Maintenance Requirements

The Project Leader is responsible for ensuring equipment and instruments are maintained according to standard operating procedures and manufacturer requirements.  In preparing for a sampling event, equipment will be inspected and tested by the sampler prior to its intended use.  A maintenance log will be maintained by the Project Leader for all mechanical and electronic equipment. Any equipment that does not meet the requirements necessary for producing data in accordance with the data quality objectives of specific projects will not be used for sample collection or analysis.  Additional equipment (non-mechanical and non-electrical), including buckets, rope, thermometers etc. should be maintained according to the standard operating procedure.

Table 15a - Equipment for Project

| Equipment Type | Manufacturer | Inspection Frequency | Type of Inspection |
| --- | --- | --- | --- |
| DO Meter |  |  |  |
| Multiprobe model: |  |  |  |
| pH meter model: |  |  |  |
| Conductivity meter model: |  |  |  |
| Turbidity meter model: |  |  |  |
| GPS Unit |  |  |  |

16.  Instrument Calibration and Frequency

The Project Leader will ensure that all field instruments are checked for good working order prior to the day of sample collection, preferably at least 24 hours prior to sampling.  On the day of sample collection, or on a routine schedule as defined below, equipment will be calibrated and checked for accuracy before any samples are collected in accordance with the standard operating procedures. The recalibration of meters will be verified by recording each meter’s reading of a standard used (or against a calibration instrument).  If the amount of drift in instrument readings is not acceptable, data will be flagged as suspect. Calibration checks and readings of standards will be recorded on field sheets or another form set up for that purpose.  All documentation regarding instrument calibration will be maintained by the Project Leader or their designated individual.

Table 16a - Equipment Calibration

| Equipment Type | Calibration Frequency | Standard or Calibration Instrument Used |
| --- | --- | --- |

17.  Inspection/Acceptance Requirements

The Project Leader will ensure that all equipment, instruments and supplies are clean and maintained according to the standards and conditions required to meet project objectives.  Sample containers will be of the appropriate size, pre-cleaned for the parameter for which the sample will be analyzed, and supplied by VAEL.  Appropriate containers must be used.  Bottles not supplied by the VAEL are considered suspect and samples will be rejected, unless lot certification of bottles is provided along with the sample submission. Other materials, such as nets, gloves, rinse bottles, sampling apparatus, buckets, line, etc., will be kept clean and stored properly so as to prevent contamination that interferes with producing samples and analytical results that meet project objectives.

18.  Data Acquisition Requirements

External data (data that is not generated by the project but is to be used as part of the project e.g., meteorological data, flow data) will be used in accordance with the objectives stated in Section 6B of this QAPP, and should have sufficient documentation that it is at least equivalent to the data quality generated as part of this project (see Section 7).

19.  Data Management

The generation of accurate data with accompanying documentation, such as field sheets and quality control sample results, is the responsibility of the individual Project Leaders.  On a daily basis, field data sheets are inspected and signed by the people performing the sampling before leaving a site or completing a sampling “run.”  Field sheets are given to the Field Leader after the sampling event for review.  Within 72 hours, the Leader will contact any samplers whose field sheets contain significant errors or omissions.

The lab manager reviews analytical results, and identifies questionable data with regard to results or documentation, as described in the LaRosa Laboratory QA Plan.  The Project Leader and the person responsible for QA review all field and lab data to determine usability in the project.

All environmental data generated by projects funded by VTDEC under this project will be submitted to the VTDEC in a commonly used format (such as Microsoft EXCEL© or ACCESS©). After additional QA review, this data will be stored in WQX and later uploaded to STORET, the national water quality data storage system.

The data generated under the laboratory services grants project is the joint property of the VTDEC and the project leads.

20.  Assessment and Response Actions

For each project funded, there will be an on-site visit by the Project Leader or Quality Assurance Coordinator to observe field sampling and field analysis procedures. Generally, this will be done near the beginning of the project.  This is in addition to training procedures described in Section 8.  A written checklist should be used for the assessments, maintained by the Project Leader, and copies will be provided with the data report.  The Project Leader and QA Coordinator will determine if field work follows the written procedures or if there needs to be corrections by additional training or revising protocols.  Please refer to Section 22 for additional evaluations and response actions regarding data evaluations.

21.  Reports

Written final project reports will be submitted to the VTDEC for all funded projects.  These need not be excessively long, but should document data results, quality assurance findings, and any specific local actions suggested by the data results.  The reports may vary in content according to the type of project and the expected uses of the information.  VTDEC strongly encourages project leaders to plan at least one presentation of their project and its results to the local community.

In addition to a written report, data and metadata (information about the data) will be provided as described in Section 19 above.

22.  Data Review, Validation, and Verification

All data are reviewed by the individual Project Leader, QA Coordinator, and Data Management Coordinator to determine if data meet QAPP requirements.

Data Analysis QC Checks will include:

- Data entry checks by a second person

- Calculation of measures of data quality.

To validate and verify project data, the project QA Coordinator will compare computer entries to field or laboratory data sheets; look for data gaps and unexpected, or nonsensical results; inspect field forms and information; review field quality control checks and resulting information; and review graphs, tables and other presentations of data, as needed.  Graphing data results with time, by parameter, is a useful way to observe problem data points.

Errors in data entry will be corrected. Data that are outside the expected range will be flagged for further review or rejected.  A second field sample and/or laboratory aliquot will be taken, if possible, to verify the condition and a determination of necessary corrections, if any, will be made. VTDEC should be contacted if assistance is needed to identify sources of errors.   Problems with data quality will be discussed in the draft and final reports to the VTDEC.  The Percent Complete table presented in Section 7 will be filled in and included with the data report.

23.  Validation and Verification Methods

The following simple measures of data quality should be calculated, and included in the final report:

1) To screen for contamination, the average blank concentration, by parameter, should be calculated.  This average value should be as close as practical to the Reporting Limit listed in Table 7b.

2) To assess the precision of results, the “Mean Relative Percent Difference” between field duplicate samples should be calculated.  The average RPD should be less than or equal to the Estimated Precision listed in Table 7b.  This simple measure is calculated as follows:

RPDfield duplicate pair 1 = absolute value (sample1 - sample2) / average (sample1 and sample2);

and,

The Mean RPD for “n” duplicate pair = average (RPDpair 1 + RPD pair 2 + ... + RPD pair n)

24. Reconciliation with Project Quality Objectives (PQOs)

As indicated above, mean blank concentrations and mean relative percent differences will be compared to data quality objectives established in Table 7b.

# Text boxes

INSTRUCTIONS: Please fill in the spaces below with appropriate information for your project and organization.  Collection of samples for this project must not take place until the QAPP is delivered to VTDEC for signature.

INSTRUCTIONS: Change page numbers and appendices as needed for your project.  Insert information for any pages of additional information you attach (e.g., maps, manuals, written procedures, etc.)

INSTRUCTIONS: please fill in the following section with the names and contact information (i.e.,address, phone #, email address) of those involved with your project who should be familiar with your QAPP.  This should include the project leader, field/sampling leader and quality assurance leader.

INSTRUCTIONS: Please fill in the name and affiliation (if not from your organization) of the person that best corresponds to the title and description in the left column.  Note that one person may have more than one responsibility and may be listed more than once, however, the person responsible for QA should not be the project leader, if possible.  If you are not using a laboratory, put an N/A (Not Applicable) in the name space.  Add other key people as needed

Instructions: Fill in the following table with the correct dates for your project.  If your project does not include any of the listed activities, note why. If there are activities not listed, add them to the table. If you have already completed a timetable you may attach it in lieu of this one.

Instructions: For the following sections (B, C, D), which address data representativeness, comparability and completeness, the VTDEC maintains a minimum goal of 80%.  On rare occasions a project requires higher goals and this may be a point of discussion during the review of your QAPP. If you think your project might be unable to meet the minimum goal, please provide the information in the lines provided below each element.

Instructions: Add any additional information on documentation and records, if applicable.

Instructions: There are a number of acceptable sampling designs depending on the purpose of your project.  Please review the descriptive information, adding project specific information as needed in the space provided.  On the following table check the appropriate box(es) and insert the numbers that reflect your project sampling design.  Submit an additional map or diagram which locates all sampling sites and important landmarks, if map under #6 does not show this. If you have this information prepared in another format, it can be substituted here.

Instructions: Please fill in and modify the following table as appropriate for your project.  If you have a separate summary of this information, you can attach it in place of this table.

INSTRUCTIONS: If your sampling methods are listed in Appendix A, please list the specific protocols you are using in the table above.  If your sampling protocol is different from the descriptions in Sections 7 and 10 or the examples in Appendix A, please attach your protocol(s) to this QAPP.

Check off the appropriate parameters in the table below.

Instructions: Please attach copy of your project-specific field sampling form here.

Instructions: For sections A, B, and C, check only those that are applicable to your project.  The goal for quality control checks is 10% replication and blank analysis.  Please note if your goal varies from this.

Instructions: Please complete the table below.